## Slide 1: An independent comparison of VARSKIN v1.2 and Geant4 Monte Carlo code for skin dose calculations
Gregory James
Head of Nuclear Medicine Physics
Royal Stoke University Hospital
United Kingdom (NHS)

## Slide 2: Aim
VARSKIN is becoming more commonplace for skin dose calculations in nuclear medicine departments in the UK.
However, assumptions and approximations made in the software may cause inaccuracies (e.g. cylinders used for droplet shapes).
The aim of this project was to independently compare the VARSKIN (v1.2) skin dose module against Geant4 Monte Carlo code for typical skin contamination scenarios.

## Slide 3
vs.

## Slide 5: Method – Geometry Considerations
Four basal depths:
70 µm
140 µm
220 µm
370 µm
Covens et al. JRP 2013
Three glove layers:
No gloves
0.1mm thickness
0.2mm thickness
Density, ρ = 0.92 g/cm3
Four geometries:
Disk model
10 µl droplet
30 µl droplet
50 µl droplet
James et al. NMC 2023
28 radionuclides:
99mTc
18F
90Y
131I
and 24 others!

## Slide 6: Geant4 Geometry
Droplet
(10, 30, 50 μl)
Sensitive
Volume
A = 1 cm2
Skin (ICRP)
10 μm
70 μm
140 μm
220 μm
370 μm
2 cm
4 cm
Skin Surface
0 mm
0.1 mm
0.2 mm
Glove
Layer
Target: >10,000 hits to achieve <1% random error
ρ = 1.1 g/cm3

## Slide 7: VARSKIN Geometry
Activity set to 1 MBq to derive instantaneous dose rate constant [mSv/MBq/hr]
Dose averaging area set to 1 cm2
Diameter set to 11.28 mm for 1 cm2

## Slide 8
Nuclear Medicine Communications 2023, 44:366–374

## Slide 9: Initial Cylinder Approximations
Cylinder 1: 99% of the height
Cylinder 4: 99% of the diameter
Cylinder 2: ⅓ of the height difference
Cylinder 3: ⅔ of the height difference

## Slide 10: Final Cylinder Recommendations
Volume | Diameter | Height
10 μl | 2.52 mm | 2.01 mm
30 μl | 4.04 mm | 2.35 mm
50 μl | 5.02 mm | 2.53 mm

## Slide 11: VARSKIN Geometry
Nuclear Medicine Communications 2023, 44:366-374
Volume | Diameter | Height
10 μl | 2.52 mm | 2.01 mm
30 μl | 4.04 mm | 2.35 mm
50 μl | 5.02 mm | 2.53 mm

## Slide 12: Results
 |  |  |  | Dose Rate (mSv/MBq/hr) |  | 
Radionuclide | Geometry | Glove Layer | Basal Depth | Geant4 | VARSKIN | Difference
99mTc | 1cm2 Disk | - | Hp(0.07) | 200 | 206 | +3%
 |  |  | Hp(0.14) | 40.9 | 55.6 | +36%
 |  |  | Hp(0.22) | 3.45 | 4.11 | +19%
 |  |  | Hp(0.37) | 2.69 | 2.76 | +3%
 |  | 0.1 mm | Hp(0.07) | 20.8 | 23.7 | +14%
 |  |  | Hp(0.14) | 3.35 | 3.66 | +9%
 |  |  | Hp(0.22) | 2.86 | 3.01 | +5%
 |  |  | Hp(0.37) | 2.46 | 2.49 | +1%
 |  | 0.2 mm | Hp(0.07) | 2.85 | 3.29 | +16%
 |  |  | Hp(0.14) | 2.75 | 2.95 | +7%
 |  |  | Hp(0.22) | 2.64 | 2.64 | +0%
 |  |  | Hp(0.37) | 2.35 | 2.28 | -3%
1344 rows of data!
(28 radionuclides × 48 geometries)
!
!
!
!

## Slide 13
VARSKIN
Geant4

## Slide 18
Linear Scale
Log Scale

## Slide 19: Results – Relative Difference
VARSKIN higher than Geant4
VARSKIN lower than Geant4

## Slide 20: Results – Relative Difference
VARSKIN higher than Geant4
VARSKIN lower than Geant4

## Slide 21: Results – Absolute Difference
VARSKIN higher than Geant4
VARSKIN lower than Geant4

## Slide 22: Results – All Data (N=1248)
 | Relative Difference (%) | Absolute Difference (mSv/MBq/hr)
Median | +7% | +0.8
95% CI | -30% to +270% | -260 to +40
Skewness | 7.1 | -2.0

## Slide 23: Literature Review
Factor Difference for Point Source |  | Factor Difference for Cylinder Source | 
No absorber | 0.96 | No absorber | 14
0.2mm Plastic | 1.8 |  | 
1.0mm Plastic | 9.3 | 1.0mm Plastic | 29
2mm Glass | 30 | 2mm Glass | 41
10mm Glass | 22 |  | 
51Cr

## Slide 24: Conclusion
Reasonable agreement between VARSKIN and Geant4 for most radionuclides.
Large relative differences observed for low dose radionuclides ( 57Co and 51Cr)
Large absolute differences observed for high dose radionuclides (68Ga, 42K, 82Rb and 90Y)
VARSKIN tends to underestimate the dose from high energy β-emitters and overestimate the dose from both low energy β-emitters and photon emitters.
This work indicates that VARSKIN v1.2 is a reliable tool for skin dose assessment for a variety of geometries and radionuclides.

## Slide 25: Thank you!
Are there any questions?